## Requirements
Requirements | Value
Positional accuracy | 0.1 degrees
Positional repeatability | 0.05 degrees
Resolution | 0.01 degrees
Motion range | Central axis to rotate from -4.5 degrees through to 45.5 degrees (0 degrees is perpendicular to the sample position)
Max speed | 0.1 deg/s
Travel time | The assembly shall travel through the range of motion within 15 minutes
Switching Positions | Switches S1 to S5 according to figure below
Switching Accuracy | 0.1 degrees

## Vaccum Tank Summary
Test: | Description: | Value | Status: | Comment:
1 | General Inspection
1.1 | Mechanical
1.1.1 | Observations
1.1.2 | Observations
1.1.3 | Observations
1.2 | Electrical
1.2.1 | Observations
1.2.2 | Grounding
1.2.3 | Motor Phase A
1.2.4 | Motor Phase B
1.2.5 | Low Limit Switch
1.2.6 | High Limit Switch
1.2.7 | Anti-Collision Switch
1.2.8 | Low Kill Switch
1.2.8 | High Kill Switch
2 | Initial Motion Test
2.1 | Observations
2.1 | Observations
Motion Performance
3 | Range and switch performance
3.1 | Range (hard stop to hardstop)
3.2 | Low Limit
3.3 | Low Kill
3.4 | Anti -Collision
3.5 | High Limit
3.6 | High Kill
4 | High speed test
5 | Accuracy
6 | Bidirectional repeatability
· Switch activation test – As per FAT
· Motion Range Test – As per FAT
· Travel with high speed – As per FAT but with feedback to produce speed curve (S2 to S4 is ok for me)
· Unidirectional Positioning Test – As per FAT
· Bidirectional Repeatability Test – Lighter version of FAT
For the last test, we should use points along the full range of motion but we don’t need to be as rigorous as the FAT was. I think the 5 positions suggested by AVS is fine but maybe we can go to 6 measurements per position? I am open to further suggestions.

## Initial motion test
Test |  | 2
Test description |  | Initial motion tests
Preparations |  | Check mechanical and electrical connections
Tests | Comments
2.1 | Posital encoder look damaged
Anti collison switch needed adjustment
Would have been safetrv to implement the switch in a different way (engaged when allowed to move, same with limits)
50deg på 15minuter |  |  |  |  | Posital value at SAM zero |  | 803488 | (dir wrong)
Gear ratio deg_m/deg |  | 12381.9095
Corerction factor |  | 0.99697
Deg | Deg_m | Deg_m with Corr
-10 | -123819.095 | -123443.8253
-5 | -61909.5475 | -61721.9127 |  | 360
0 | 0 | 0 | Perpedicular to sample | 0.029075 | 1.7445
5 | 61909.5475 | 61721.9127
10 | 123819.095 | 123443.8253
15 | 185728.6425 | 185165.738
20 | 247638.19 | 246887.6507
25 | 309547.7375 | 308609.5633
30 | 371457.285 | 370331.476 |  |  |  | Move 5 deg:
35 | 433366.8325 | 432053.3886 |  |  |  | start | 11183
40 | 495276.38 | 493775.3013 |  |  |  | 5 deg | 62000
45 | 557185.9275 | 555497.214 |  |  |  | Target in deg_motor | 73183
48.5 | 600522.6107 | 598702.5528

## Motion range, switch perf
Test |  | 3
Test description |  | Motion range and switch performance test
Engage and disenage each switch 5 times.
Preparations |  | Preparations go to Low hardstop
Test number |  | Laser Sacnner | Posital encoder | Open loop counter |  |  |  | Tests | Comments
3.1 | Low hard stop | [6 empty] | 3
| Switch | Laser Sacnner |  | Posital encoder |  | Open loop counter
3.2 | L1 | Engage | Disengage | Engage | Disengage | Engage | Disengage
1
2
3
4
5
| Switch | Laser Sacnner |  | Posital encoder |  | Open loop counter
3.3 | L2 | Engage | Disengage | Engage | Disengage | Engage | Disengage
1
2
3
4
5
| Switch | Laser Sacnner |  | Posital encoder |  | Open loop counter
3.4 | L3 | Engage | Disengage | Engage | Disengage | Engage | Disengage
1
2
3
4
5
| Switch | Laser Sacnner |  | Posital encoder |  | Open loop counter
3.5 | L4 | Engage | Disengage | Engage | Disengage | Engage | Disengage
1
2
3
4
5
| Switch | Laser Sacnner |  | Posital encoder |  | Open loop counter
3.6 | L5 | Engage | Disengage | Engage | Disengage | Engage | Disengage
1
2
3
4
5
|  | Laser Sacnner | Posital encoder | Open loop counter
3.1 | High hard stop

## High Speed
Test |  | 4
Test description |  | High speed test
Run for low limit switch to high limit switch in high speed.
Record speed and position
Tests | Comments
4

## Accuracy
Test | 5
Test description |  | Accuarcy | Move to 10 positions and record
Motor deg | Position | Laser tracker | Open loop position | Posital Encoder | Tests | Comments
| -5 |  |  |  | 5
0
5
10
15
20
25
30
35
40

## Repeatabilty
Test |  | 6
Test description |  | Bidirectional repeatability
Approach 6 different positions form both directions. Repeat 6 times
Positions to appraoch |  | -5 | 0 | 10 | 20 | 30 | 40
Position | -5deg | Approach 6 different positions form both directions. Repeat 6 times -5: from below
Test | 1 | 2 | 3 | 4 | 5 | 6 |  | Tests | Comments
Laser scanner position | [7 empty] | 6
Open loop counter
Posital encoder
Position | -5deg | Approach 6 different positions form both directions. Repeat 6 times -5: from above
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position
Open loop counter
Posital encoder
Position | 0deg | Approach 6 different positions form both directions. Repeat 6 times -5: from below
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position
Open loop counter
Posital encoder
Position | 0deg | Approach 6 different positions form both directions. Repeat 6 times -5: from above
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position
Open loop counter
Posital encoder
Position | 10deg | Approach 6 different positions form both directions. Repeat 6 times -5: from below
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position
Open loop counter
Posital encoder
Position | 10deg | Approach 6 different positions form both directions. Repeat 6 times -5: from above
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position
Open loop counter
Posital encoder
Position | 20deg | Approach 6 different positions form both directions. Repeat 6 times -5: from below
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position
Open loop counter
Posital encoder
Position | 20deg | Approach 6 different positions form both directions. Repeat 6 times -5: from above
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position
Open loop counter
Posital encoder
Position | 30deg | Approach 6 different positions form both directions. Repeat 6 times -5: from below
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position
Open loop counter
Posital encoder
Position | 30deg | Approach 6 different positions form both directions. Repeat 6 times -5: from above
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position
Open loop counter
Posital encoder
Position | 40deg | Approach 6 different positions form both directions. Repeat 6 times -5: from below
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position
Open loop counter
Posital encoder
Position | 40deg | Approach 6 different positions form both directions. Repeat 6 times -5: from above
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position
Open loop counter
Posital encoder