## Requirements
Requirements | Value
Positional accuracy | 0.1 degrees
Positional repeatability | 0.05 degrees
Resolution | 0.01 degrees
Motion range | Central axis to rotate from -4.5 degrees through to 45.5 degrees (0 degrees is perpendicular to the sample position)
Max speed | 0.1 deg/s
Travel time | The assembly shall travel through the range of motion within 15 minutes
Switching Positions | Switches S1 to S5 according to figure below
Switching Accuracy | 0.1 degrees

## Vaccum Tank Summary
Test: | Description: | Value | Status: | Comment:
1 | General Inspection
1.1 | Mechanical
1.1.1 | Observations |  | OK | Dent on posital encoder housing. Encoder was concluded to work correct.
1.2 | Electrical
1.2.1 | Observations
1.2.2 | Grounding |  | OK
1.2.3 | Motor Phase A |  | OK
1.2.4 | Motor Phase B |  | OK
1.2.5 | Low Limit Switch |  | OK
1.2.6 | High Limit Switch |  | OK
1.2.7 | Anti-Collision Switch |  | OK | Needed adjustment
1.2.8 | Low Kill Switch |  | OK
1.2.8 | High Kill Switch |  | OK
2 | Initial Motion Test
2.1 | Gera ratio posital encoder [] | -4.4525e-05 | OK
2.1 | Gera ratio open loop [] | 8.1062e-05 | OK
2.1 | Backlash [deg] | 0.01 | OK | Based on one measurement at -10 deg
Motion Performance
3 | Range and switch performance |  |  |  |  | Gearratio | Offset | Accuracy
3.1 | Range (hard stop to hardstop) |  | Not ok | Could not reach upper hard limit | Openloop | 8.1062e-05 | -0.0166 | 0.0039
3.2 | Low Kill Engage [deg] | -4.606 | OK |  | Posital | -4.4525e-05 | 30.7262 | 0.0067
3.2 | Low Kill Disengage [deg] | -4.5738 | OK |  | Posital | -4.4525e-05 | 30.7262 | 0.0067
3.3 | Low Limit Engage [deg] | -4.281 | OK
3.3 | Low Limit Disengage [deg] | -4.2535 | OK
3.4 | Anti -Collision Engage [deg] | 32.8626 | OK
3.4 | Anti -Collision Disengage [deg] | 32.8311 | OK
3.5 | High Limit Engage [deg] | 44.965 | OK
3.5 | High Limit Disengage [deg] | 44.9225 | OK
3.6 | High Kill Engage [deg] | 45.35 | OK
3.6 | High Kill Disengage [deg] | 45.3211 | OK
5 | Accuracy open loop [deg] | 0.0039 | OK | Zero position was treated as a outlier and therefore excluded.
5 | Accuracy Posital encoder [deg] | 0.0067 | OK | Zero position was treated as a outlier and therefore excluded.
6 | Bidirectional repeatability [deg] | 0.0088 | OK | Zero position was treated as a outlier and therefore excluded.
2 | Initial Motion Test
2.1 | Gera ratio posital encoder [] | -4.4525e-05 | OK | Seems to have changed slightly |  | Gearratio | Offset | Accuracy
2.1 | Gera ratio open loop [] | 8.1062e-05 | OK |  | Openloop | 8.1062e-05 | 8.1062e-05 | 0.01466
2.1 | Backlash [deg] | 0.02 | Check | Based on highe speed test at 45 deg. Need more data to conclude. | Posital | -4.4525e-05 | 30.7262 | 0.0037821
Motion Performance
4 | High speed test |  | OK
5.1 | Accuracy open loop [deg] | 0.01466 | Check | Open loop accuracy seems to be worse after collision. But data is based on only a few measurements.
5.2 | Accuracy Posital encoder[deg] | 0.0037821 | OK | Even better than before collision but fewer measurements
6 | Bidirectional repeatability [deg] |  | Check | Not measured after collision |  | Engage |  | Disengage
[5 empty] | Switch | Position [deg] | Range [deg] | Position [deg] | Range [deg]
[5 empty] | Low kill | -4.606 | 0.00071218 | -4.5738 | 0.00031158
[5 empty] | Low limit | -4.281 | 0.00084572 | -4.2535 | 0.00075669
[5 empty] | Anti collision | 32.8626 | 0.00031158 | 32.8311 | 0.00035609
[5 empty] | High limit | 44.965 | 0.0004006 | 44.9225 | 0.0004006
[5 empty] | High kill | 45.35 | 0.00031158 | 45.3211 | 0.00035609

## Initial motion test
| Test |  | 2
| Test description |  | Initial motion tests
| Preparations |  | Check mechanical and electrical connections
| Tests | Comments
| 2.1 | Posital encoder look damaged
Anti collison switch needed adjustment
Would have been safetrv to implement the switch in a different way (engaged when allowed to move, same with limits)
[12 empty] | 12343.1669 | 296236.0065
[7 empty] | 44.94 | 45 | 1.0013 |  | 296236.0065
| 50deg på 15minuter | [9 empty] | Posital value at SAM zero |  | 803488 | (dir wrong)
|  |  | AVS | ESS | 8.1016e-05
| Gear ratio deg_m/deg |  | 12381.9095 | 12326.7094 |  | 8.0763e-05
| Corerction factor |  | 0.99697 | 1.0013
| Deg | Deg_m (AVS) | Deg_m with Corr (AVS) | Encoder | Laser | open loop
| -10 | -123819.095 | -123443.8253 |  |  | -123431.6694
| -5 | -61909.5475 | -61721.9127 |  |  | -61715.8347 |  |  |  | 360
Perpedicular to sample | 0 | 0 | 0 |  |  | 0 |  |  |  | 0.029075 | 1.7445
| 5 | 61909.5475 | 61721.9127 |  |  | 61715.8347
| 10 | 123819.095 | 123443.8253 |  |  | 123431.6694
| 15 | 185728.6425 | 185165.738 |  |  | 185147.504
| 20 | 247638.19 | 246887.6507 |  |  | 246863.3387
| 25 | 309547.7375 | 308609.5633 |  |  | 308579.1734
| 30 | 371457.285 | 370331.476 |  |  | 370295.0081 | [5 empty] | Move 5 deg:
| 35 | 433366.8325 | 432053.3886 |  |  | 432010.8428 | 435310.9037 |  |  |  |  | start | 11183
| 40 | 495276.38 | 493775.3013 |  |  | 493726.6774 | [5 empty] | 5 deg | 62000
| 45 | 557185.9275 | 555497.214 |  |  | 555442.5121 | 558742.573 |  |  |  |  | Target in deg_motor | 73183
| 48.5 | 600522.6107 | 598702.5528 |  |  | 598643.5964
|  | High speed test in 0.1deg/s |  | 1234.3167 | motordeg/s
1.1287e-05

## Accuracy
Test | 5
Test description |  | Accuarcy | Move to 5 positions and record
After Collision
Motor deg | Position | Open loop | Posital Encoder | Diff OL Posital |  | Gearratio | Offset |  | Accuracy | Tests | Comments
| 0 | 0 | 0 | 0 | Openloop | 8.1016e-05 | -3.4179e-06 | Max value= | 0.0064535
| 10 | 10 | 10.0065 | 0.0064535 | Posital | 4.4503e-05 | 30.7613 | Average= | 0.0035338 | 5
| 20 | 20 | 20.0065 | 0.0064535
| 30 | 30 | 29.9959 | 0.0041371 | Gear ratio DIFF (before vs after collision)
| 40 | 40 | 40.0006 | 0.00062475 | Open loop | 2.4764e-08
[5 empty] | posital | 8.2131e-09
[6 empty] | Posital | OL
[5 empty] | Value initial Zero: | 691207 | 0
| Position | Open loop raw | Posital Encoder Raw |  | Value Zero after test | 690480 | 0.04921875
| 0 | -0.042188 | 691215
| 10 | 123431.625 | 466367
| 20 | 246863.2992 | 241664 |  |  |  |  | Before coll
| 30 | 370294.9594 | 17199 |  | 123431.6742 | 10 | 8.1016e-05 | 8.1041e-05
| 40 | 493726.6336 | 2147276036 |  | 224703 |  | 4.4503e-05 | 4.4511e-05
|  |  | 2147483647 |  |  |  | 30.7613
35
45
hardstop high

## Repeatability
Test |  | 6
Test description |  | Bidirectional repeatability
Approach 6 different positions form both directions, Repeat 6 times
Positions to approach |  | -5 | 0 | 10 | 20 | 30 | 40
[9 empty] | Tests | Comments
Position | 0deg | Approach 6 different positions form both directions, Repeat 6 times 0: from below; Approach 6 different positions form both directions, Repeat 6 times 30: After Collision
Test | Open loop raw | Open loop counter | Posital encoder raw | Posital encoder |  |  | Gearratio | Offset
1 | -0.042188 | 0 | 691215 | 0 |  | Openloop | 8.1016e-05 | -3.4179e-06
2 | -0.042188 | 0 | 691212 | 0.00013351 |  | Posital | 4.4503e-05 | 30.7613
3 | -0.049219 | Approach 6 different positions form both directions, Repeat 6 times -5: -5.6965e-07; Approach 6 different positions form both directions, Repeat 6 times 0: 691212; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00013351
4 | -0.049219 | Approach 6 different positions form both directions, Repeat 6 times -5: -5.6965e-07; Approach 6 different positions form both directions, Repeat 6 times 0: 691212; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00013351
Range |  | Approach 6 different positions form both directions, Repeat 6 times -5: 5.6965e-07; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00013351
STD |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0.0040595; Approach 6 different positions form both directions, Repeat 6 times 10: 6.6755e-05; Approach 6 different positions form both directions, Repeat 6 times 40: System repeatabilty
Repeatabilty | 0.00013351 | Approach 6 different positions form both directions, Repeat 6 times 30: Max. value; Approach 6 different positions form both directions, Repeat 6 times 40: 0.00093457
[6 empty] | Average | 0.00060079
Position | 0deg | Approach 6 different positions form both directions, Repeat 6 times 0: from above
Test | Open loop raw | Approach 6 different positions form both directions, Repeat 6 times -5: Open loop counter; Approach 6 different positions form both directions, Repeat 6 times 0: Posital encoder raw; Approach 6 different positions form both directions, Repeat 6 times 10: Posital encoder
1 | 0.049219 | Approach 6 different positions form both directions, Repeat 6 times -5: 7.4054e-06; Approach 6 different positions form both directions, Repeat 6 times 0: 690520; Approach 6 different positions form both directions, Repeat 6 times 10: 0.03093
2 | 0.042188 | Approach 6 different positions form both directions, Repeat 6 times -5: 6.8358e-06; Approach 6 different positions form both directions, Repeat 6 times 0: 690512; Approach 6 different positions form both directions, Repeat 6 times 10: 0.031286
3 | 0.042188 | Approach 6 different positions form both directions, Repeat 6 times -5: 6.8358e-06; Approach 6 different positions form both directions, Repeat 6 times 0: 690511; Approach 6 different positions form both directions, Repeat 6 times 10: 0.03133
4 | 0.042188 | Approach 6 different positions form both directions, Repeat 6 times -5: 6.8358e-06; Approach 6 different positions form both directions, Repeat 6 times 0: 690507; Approach 6 different positions form both directions, Repeat 6 times 10: 0.031508
Range |  | Approach 6 different positions form both directions, Repeat 6 times -5: 5.6965e-07; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00057854
STD |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0.0035156; Approach 6 different positions form both directions, Repeat 6 times 10: 0.0002424
Repeatabilty | 0.00057854
Position | 10deg | Approach 6 different positions form both directions, Repeat 6 times 0: from below
Test | Open loop raw | Approach 6 different positions form both directions, Repeat 6 times -5: Open loop counter; Approach 6 different positions form both directions, Repeat 6 times 0: Posital encoder raw; Approach 6 different positions form both directions, Repeat 6 times 10: Posital encoder
1 | 123431.625 | Approach 6 different positions form both directions, Repeat 6 times -5: 10; Approach 6 different positions form both directions, Repeat 6 times 0: 466367; Approach 6 different positions form both directions, Repeat 6 times 10: 10.0065
2 | 123431.625 | Approach 6 different positions form both directions, Repeat 6 times -5: 10; Approach 6 different positions form both directions, Repeat 6 times 0: 466352; Approach 6 different positions form both directions, Repeat 6 times 10: 10.0071
3 | 123431.625 | Approach 6 different positions form both directions, Repeat 6 times -5: 10; Approach 6 different positions form both directions, Repeat 6 times 0: 466359; Approach 6 different positions form both directions, Repeat 6 times 10: 10.0068
4 | 123431.625 | Approach 6 different positions form both directions, Repeat 6 times -5: 10; Approach 6 different positions form both directions, Repeat 6 times 0: 466363; Approach 6 different positions form both directions, Repeat 6 times 10: 10.0066
Range |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00066755
STD |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00028467
Repeatabilty | 0.00066755
Position | 10deg | Approach 6 different positions form both directions, Repeat 6 times 0: from above
Test | Open loop raw | Approach 6 different positions form both directions, Repeat 6 times -5: Open loop counter; Approach 6 different positions form both directions, Repeat 6 times 0: Posital encoder raw; Approach 6 different positions form both directions, Repeat 6 times 10: Posital encoder
1 | 123431.7094 | Approach 6 different positions form both directions, Repeat 6 times -5: 10; Approach 6 different positions form both directions, Repeat 6 times 0: 465771; Approach 6 different positions form both directions, Repeat 6 times 10: 10.033
2 | 123431.7094 | Approach 6 different positions form both directions, Repeat 6 times -5: 10; Approach 6 different positions form both directions, Repeat 6 times 0: 465767; Approach 6 different positions form both directions, Repeat 6 times 10: 10.0332
3 | 123431.7094 | Approach 6 different positions form both directions, Repeat 6 times -5: 10; Approach 6 different positions form both directions, Repeat 6 times 0: 465763; Approach 6 different positions form both directions, Repeat 6 times 10: 10.0333
4 | 123431.7164 | Approach 6 different positions form both directions, Repeat 6 times -5: 10; Approach 6 different positions form both directions, Repeat 6 times 0: 465759; Approach 6 different positions form both directions, Repeat 6 times 10: 10.0335
Range |  | Approach 6 different positions form both directions, Repeat 6 times -5: 5.6712e-07; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00053404
STD |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0.0035; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00022981
Repeatabilty | 0.00053404
Position | 20deg | Approach 6 different positions form both directions, Repeat 6 times 0: from below
Test | Open loop raw | Approach 6 different positions form both directions, Repeat 6 times -5: Open loop counter; Approach 6 different positions form both directions, Repeat 6 times 0: Posital encoder raw; Approach 6 different positions form both directions, Repeat 6 times 10: Posital encoder
1 | 246863.2992 | Approach 6 different positions form both directions, Repeat 6 times -5: 20; Approach 6 different positions form both directions, Repeat 6 times 0: 241664; Approach 6 different positions form both directions, Repeat 6 times 10: 20.0065
2 | 246863.2922 | Approach 6 different positions form both directions, Repeat 6 times -5: 20; Approach 6 different positions form both directions, Repeat 6 times 0: 241656; Approach 6 different positions form both directions, Repeat 6 times 10: 20.0068
3 | 246863.2992 | Approach 6 different positions form both directions, Repeat 6 times -5: 20; Approach 6 different positions form both directions, Repeat 6 times 0: 241663; Approach 6 different positions form both directions, Repeat 6 times 10: 20.0065
4 | 246863.2992 | Approach 6 different positions form both directions, Repeat 6 times -5: 20; Approach 6 different positions form both directions, Repeat 6 times 0: 241668; Approach 6 different positions form both directions, Repeat 6 times 10: 20.0063
Range |  | Approach 6 different positions form both directions, Repeat 6 times -5: 5.6712e-07; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00053404
STD |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0.0035; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00022214
Repeatabilty | 0.00053404
Position | 20deg | Approach 6 different positions form both directions, Repeat 6 times 0: from above
Test | Open loop raw | Approach 6 different positions form both directions, Repeat 6 times -5: Open loop counter; Approach 6 different positions form both directions, Repeat 6 times 0: Posital encoder raw; Approach 6 different positions form both directions, Repeat 6 times 10: Posital encoder
1 | 246863.3836 | Approach 6 different positions form both directions, Repeat 6 times -5: 20; Approach 6 different positions form both directions, Repeat 6 times 0: 241192; Approach 6 different positions form both directions, Repeat 6 times 10: 20.0275
2 | 246863.3836 | Approach 6 different positions form both directions, Repeat 6 times -5: 20; Approach 6 different positions form both directions, Repeat 6 times 0: 241188; Approach 6 different positions form both directions, Repeat 6 times 10: 20.0276
3 | 246863.3836 | Approach 6 different positions form both directions, Repeat 6 times -5: 20; Approach 6 different positions form both directions, Repeat 6 times 0: 241184; Approach 6 different positions form both directions, Repeat 6 times 10: 20.0278
4 | 246863.3836 | Approach 6 different positions form both directions, Repeat 6 times -5: 20; Approach 6 different positions form both directions, Repeat 6 times 0: 241180; Approach 6 different positions form both directions, Repeat 6 times 10: 20.028
Range |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00053404
STD |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00022981
Repeatabilty | 0.00053404
Position | 30deg | Approach 6 different positions form both directions, Repeat 6 times 0: from below
Test | Open loop raw | Approach 6 different positions form both directions, Repeat 6 times -5: Open loop counter; Approach 6 different positions form both directions, Repeat 6 times 0: Posital encoder raw; Approach 6 different positions form both directions, Repeat 6 times 10: Posital encoder
1 | 370294.9594 | Approach 6 different positions form both directions, Repeat 6 times -5: 30; Approach 6 different positions form both directions, Repeat 6 times 0: 17199; Approach 6 different positions form both directions, Repeat 6 times 10: 29.9959
2 | 370294.9594 | Approach 6 different positions form both directions, Repeat 6 times -5: 30; Approach 6 different positions form both directions, Repeat 6 times 0: 17180; Approach 6 different positions form both directions, Repeat 6 times 10: 29.9967
3 | 370294.9664 | Approach 6 different positions form both directions, Repeat 6 times -5: 30; Approach 6 different positions form both directions, Repeat 6 times 0: 17184; Approach 6 different positions form both directions, Repeat 6 times 10: 29.9965
4 | 370294.9664 | Approach 6 different positions form both directions, Repeat 6 times -5: 30; Approach 6 different positions form both directions, Repeat 6 times 0: 17184; Approach 6 different positions form both directions, Repeat 6 times 10: 29.9965
Range |  | Approach 6 different positions form both directions, Repeat 6 times -5: 5.6712e-07; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00084556
STD |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0.0040415; Approach 6 different positions form both directions, Repeat 6 times 10: 0.000373
Repeatabilty | 0.00084556
Position | 30deg | Approach 6 different positions form both directions, Repeat 6 times 0: from above
Test | Open loop raw | Approach 6 different positions form both directions, Repeat 6 times -5: Open loop counter; Approach 6 different positions form both directions, Repeat 6 times 0: Posital encoder raw; Approach 6 different positions form both directions, Repeat 6 times 10: Posital encoder
1 | 370295.0578 | Approach 6 different positions form both directions, Repeat 6 times -5: 30; Approach 6 different positions form both directions, Repeat 6 times 0: 16527; Approach 6 different positions form both directions, Repeat 6 times 10: 30.0258
2 | 370295.0508 | Approach 6 different positions form both directions, Repeat 6 times -5: 30; Approach 6 different positions form both directions, Repeat 6 times 0: 16523; Approach 6 different positions form both directions, Repeat 6 times 10: 30.0259
3 | 370295.0578 | Approach 6 different positions form both directions, Repeat 6 times -5: 30; Approach 6 different positions form both directions, Repeat 6 times 0: 16516; Approach 6 different positions form both directions, Repeat 6 times 10: 30.0263
4 | 370295.0578 | Approach 6 different positions form both directions, Repeat 6 times -5: 30; Approach 6 different positions form both directions, Repeat 6 times 0: 16511; Approach 6 different positions form both directions, Repeat 6 times 10: 30.0265
Range |  | Approach 6 different positions form both directions, Repeat 6 times -5: 5.6712e-07; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00071205
STD |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0.0035; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00031756
Repeatabilty | 0.00071205
Position | 40deg | Approach 6 different positions form both directions, Repeat 6 times 0: from below
Test | Open loop raw | Approach 6 different positions form both directions, Repeat 6 times -5: Open loop counter; Approach 6 different positions form both directions, Repeat 6 times 0: Posital encoder raw; Approach 6 different positions form both directions, Repeat 6 times 10: Posital encoder
1 | 493726.6336 | Approach 6 different positions form both directions, Repeat 6 times -5: 40; Approach 6 different positions form both directions, Repeat 6 times 0: 2147276036; Approach 6 different positions form both directions, Repeat 6 times 10: 40.0006
2 | 493726.6406 | Approach 6 different positions form both directions, Repeat 6 times -5: 40; Approach 6 different positions form both directions, Repeat 6 times 0: 2147276015; Approach 6 different positions form both directions, Repeat 6 times 10: 40.0016
3 | 493726.6336 | Approach 6 different positions form both directions, Repeat 6 times -5: 40; Approach 6 different positions form both directions, Repeat 6 times 0: 2147276023; Approach 6 different positions form both directions, Repeat 6 times 10: 40.0012
4 | 493726.6406 | Approach 6 different positions form both directions, Repeat 6 times -5: 40; Approach 6 different positions form both directions, Repeat 6 times 0: 2147276027; Approach 6 different positions form both directions, Repeat 6 times 10: 40.001
Range |  | Approach 6 different positions form both directions, Repeat 6 times -5: 5.6712e-07; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00093457
STD |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0.0040415; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00038861
Repeatabilty | 0.00093457
Position | 40deg | Approach 6 different positions form both directions, Repeat 6 times 0: from above
Test | Open loop raw | Approach 6 different positions form both directions, Repeat 6 times -5: Open loop counter; Approach 6 different positions form both directions, Repeat 6 times 0: Posital encoder raw; Approach 6 different positions form both directions, Repeat 6 times 10: Posital encoder
1 | 493726.718 | Approach 6 different positions form both directions, Repeat 6 times -5: 40; Approach 6 different positions form both directions, Repeat 6 times 0: 2147275436; Approach 6 different positions form both directions, Repeat 6 times 10: 40.0273
2 | 493726.718 | Approach 6 different positions form both directions, Repeat 6 times -5: 40; Approach 6 different positions form both directions, Repeat 6 times 0: 2147275431; Approach 6 different positions form both directions, Repeat 6 times 10: 40.0275
3 | 493726.718 | Approach 6 different positions form both directions, Repeat 6 times -5: 40; Approach 6 different positions form both directions, Repeat 6 times 0: 2147275428; Approach 6 different positions form both directions, Repeat 6 times 10: 40.0277
4 | 493726.718 | Approach 6 different positions form both directions, Repeat 6 times -5: 40; Approach 6 different positions form both directions, Repeat 6 times 0: 2147275424; Approach 6 different positions form both directions, Repeat 6 times 10: 40.0279
Range |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0; Approach 6 different positions form both directions, Repeat 6 times 10: 0.00053404
STD |  | Approach 6 different positions form both directions, Repeat 6 times -5: 0; Approach 6 different positions form both directions, Repeat 6 times 10: 0.0002251
Repeatabilty | 0.00053404