## Requirements
Requirements | Value
Positional accuracy | 0.1 degrees
Positional repeatability | 0.05 degrees
Resolution | 0.01 degrees
Motion range | Central axis to rotate from -4.5 degrees through to 45.5 degrees (0 degrees is perpendicular to the sample position)
Max speed | 0.1 deg/s
Travel time | The assembly shall travel through the range of motion within 15 minutes
Switching Positions | Switches S1 to S5 according to figure below
Switching Accuracy | 0.1 degrees

## Vaccum Tank Summary
Test: | Description: | Value | Status: | Comment:
1 | General Inspection
1.1 | Mechanical
1.1.1 | Observations
1.1.2 | Observations
1.1.3 | Observations
1.2 | Electrical
1.2.1 | Observations
1.2.2 | Grounding
1.2.3 | Motor Phase A
1.2.4 | Motor Phase B
1.2.5 | Low Limit Switch
1.2.6 | High Limit Switch
1.2.7 | Anti-Collision Switch
1.2.8 | Low Kill Switch
1.2.8 | High Kill Switch
2 | Initial Motion Test
2.1 | Observations
2.1 | Observations
Motion Performance
3 | Range and switch performance
3.1 | Range (hard stop to hardstop)
3.2 | Low Limit
3.3 | Low Kill
3.4 | Anti -Collision
3.5 | High Limit
3.6 | High Kill
4 | High speed test
5 | Accuracy
6 | Bidirectional repeatability
· Switch activation test – As per FAT
· Motion Range Test – As per FAT
· Travel with high speed – As per FAT but with feedback to produce speed curve (S2 to S4 is ok for me)
· Unidirectional Positioning Test – As per FAT
· Bidirectional Repeatability Test – Lighter version of FAT
For the last test, we should use points along the full range of motion but we don’t need to be as rigorous as the FAT was. I think the 5 positions suggested by AVS is fine but maybe we can go to 6 measurements per position? I am open to further suggestions.

## Initial motion test
| Test |  | 2
| Test description |  | Initial motion tests
| Preparations |  | Check mechanical and electrical connections
| Tests | Comments
| 2.1 | Posital encoder look damaged
Anti collison switch needed adjustment
Would have been safetrv to implement the switch in a different way (engaged when allowed to move, same with limits)
[12 empty] | 12343.1669 | 296236.0065
[7 empty] | 44.94 | 45 | 1.0013 |  | 296236.0065
| 50deg på 15minuter | [9 empty] | Posital value at SAM zero |  | 803488 | (dir wrong)
|  |  | AVS | ESS | 8.1016e-05
| Gear ratio deg_m/deg |  | 12381.9095 | 12326.7094 |  | 8.0763e-05
| Corerction factor |  | 0.99697 | 1.0013
| Deg | Deg_m (AVS) | Deg_m with Corr (AVS) | Encoder | Laser | open loop
| -10 | -123819.095 | -123443.8253 |  |  | -123431.6694
| -5 | -61909.5475 | -61721.9127 |  |  | -61715.8347 |  |  |  | 360
Perpedicular to sample | 0 | 0 | 0 |  |  | 0 |  |  |  | 0.029075 | 1.7445
| 5 | 61909.5475 | 61721.9127 |  |  | 61715.8347
| 10 | 123819.095 | 123443.8253 |  |  | 123431.6694
| 15 | 185728.6425 | 185165.738 |  |  | 185147.504
| 20 | 247638.19 | 246887.6507 |  |  | 246863.3387
| 25 | 309547.7375 | 308609.5633 |  |  | 308579.1734
| 30 | 371457.285 | 370331.476 |  |  | 370295.0081 | [5 empty] | Move 5 deg:
| 35 | 433366.8325 | 432053.3886 |  |  | 432010.8428 | 435310.9037 |  |  |  |  | start | 11183
| 40 | 495276.38 | 493775.3013 |  |  | 493726.6774 | [5 empty] | 5 deg | 62000
| 45 | 557185.9275 | 555497.214 |  |  | 555442.5121 | 558742.573 |  |  |  |  | Target in deg_motor | 73183
| 48.5 | 600522.6107 | 598702.5528 |  |  | 598643.5964
|  | High speed test in 0.1deg/s |  | 1234.3167 | motordeg/s
1.1287e-05

## Motion range, switch perf
Test |  | 3
Test description |  | Motion range and switch performance test
Engage and disenage each switch 5 times.
Preparations |  | Preparations go to Low hardstop
Test number |  | Laser Sacnner | Posital encoder | Open loop counter |  |  |  | Tests | Comments
3.1 | Low hard stop | 10.95 | 934879 | -76718.98828 |  |  |  | 3
| Switch | Laser Sacnner |  | Posital encoder |  | Open loop counter
3.2 | L1 | Engage | Disengage | Engage | Disengage | Engage | Disengage
1
2
3
4
5
| Switch | Laser Sacnner |  | Posital encoder |  | Open loop counter
3.3 | L2 | Engage | Disengage | Engage | Disengage | Engage | Disengage
1
2
3
4
5
| Switch | Laser Sacnner |  | Posital encoder |  | Open loop counter
3.4 | L3 | Engage | Disengage | Engage | Disengage | Engage | Disengage
1
2
3
4
5
| Switch | Laser Sacnner |  | Posital encoder |  | Open loop counter
3.5 | L4 | Engage | Disengage | Engage | Disengage | Engage | Disengage
1
2
3
4
5
| Switch | Laser Sacnner |  | Posital encoder |  | Open loop counter
3.6 | L5 | Engage | Disengage | Engage | Disengage | Engage | Disengage
1
2
3
4
5
|  | Laser Sacnner | Posital encoder | Open loop counter
3.1 | High hard stop

## High Speed
Test |  | 4
Test description |  | High speed test
Run for low limit switch to high limit switch in high speed.
Record speed and position
Tests | Comments
4

## Accuracy
Test | 5
Test description |  | Accuarcy | Move to 10 positions and record
BEFORE COLLISION
Motor deg | Position | Laser tracker | Open loop position | OL diff abs | Posital Encoder | Posital Diff abs | [8 empty] | Tests | Comments
| hardstop low | -10.98 | -10.9782 | 0.0017896 | -10.9744 | 0.0055828
| -5 | -5,0177 | -5.0097 | 0.0080226 | -5.0102 | 0.007499 | [8 empty] | 5
| 0 | 0.0192 | -0.0081487 | 0.027349 | -0.011749 | 0.030949 |  |  |  | Gearratio | Offset | Accuracy
| 5 | 4.9862 | 4.9934 | 0.0071788 | 4.9956 | 0.0094061 |  |  | Openloop | 8.1041e-05 | -0.0081453 | 0.027349
| 10 | 9.9888 | 9.9949 | 0.0061074 | 9.9986 | 0.0098431 |  |  | Posital | -4.4511e-05 | 30.7267 | 0.030949
| 15 | 14.9897 | 14.9964 | 0.0067355 | 15.0016 | 0.011891
| 20 | 19.9951 | 19.998 | 0.0028636 | 19.9979 | 0.002807
| 25 | 24.9998 | 24.9995 | 0.00030834 | 24.9971 | 0.0027285
| 30 | 29.9965 | 30.001 | 0.0045197 | 29.9951 | 0.0013767
| 35 | 35.0005 | 35.0025 | 0.0020484 | 34.9974 | 0.0031409
| 40 | 40.0095 | 40.0041 | 0.0054235 | 40.0049 | 0.0046082
| 45 | 45.0127 | 45.0056 | 0.0070955 | 45.0114 | 0.0013437
| hardstop high | 48.5
Over 48,5 then collision with concrete
BELOW AFTER COLLISION
Rereferenced after the wall incident
| Position | Laser tracker | Lasertracker | Open loop position |  |  |  |  | Posital Encoder
| hardstop low | 10.98 |  |  | Offset | scale OL | Deg | Diff |  | overflow | Offset | Scale | Deg | Diff
| -5 | -5,0017 |  | -58415.7305 | -61715.7914 | 8.1044e-05 | -4.9916 | -0.010123 | 802915 | 802915 | 112432 | -4.4486e-05 | -5.0045 | 0.0028048
| 0 | 0 | coming from 45 = 0,0042 (lasertracker) | 3300.0609 | 0 |  | 0 | 0 | 690483 | 690483 | 0 |  | 0
| 5 | 4.9725 |  | 65015.7961 | 61715.7352 | 8.0571e-05 | 4.9916 | -0.019072 | 578788 | 578788 | -111695 | -4.4519e-05 | 4.9717 | 0.00080006
| 10 | 9.9755 |  | 126731.6297 | 123431.5688 | 8.0818e-05 | 9.9832 | -0.0076524 | 466368 | 466368 | -224115 | -4.4511e-05 | 9.9757 | -0.00017063
15
20
25
30
| 35 | 34.9801 |  | 435310.8609 | 432010.8 | 8.097e-05 | 34.9411 | 0.039041 | 2147388488 | -95160 | -785643 | -4.4524e-05 | 34.9701 | 0.010036
40
| 45 | 44.989 |  | 558742.5281 | 555442.4672 | 8.0997e-05 | 44.9242 | 0.064781 | 2147163532 | -320116 | -1010599 | -4.4517e-05 | 44.9832 | 0.0058317
hardstop high
[5 empty] | Average scale: | 8.088e-05 |  |  |  |  | Average scale: | -4.4511e-05
[5 empty] | Offset | -3300.0609 |  |  |  |  | Offset | -690483
After collision
| Position | Laser tracker | Open loop position | OL diff abs | Posital Encoder | Posital Diff abs
| hardstop low | 10.98
| -5 | -5.0017 | -5.0112 | 0.0094989 | -5.0043 | 0.0026321
| 0 | 0 | -0.011927 | 0.011927 | 0.00093084 | 0.00093084 |  |  |  | Gearratio | Offset | Accuracy
| 5 | 4.9725 | 4.9873 | 0.014834 | 4.9734 | 0.00088387 |  |  | Openloop | 8.1005e-05 | -0.27925 | 0.014834
| 10 | 9.9755 | 9.9866 | 0.011106 | 9.9781 | 0.0026125 |  |  | Posital | -4.4518e-05 | 30.7399 | 0.0038084
| 15 |  |  | 0 |  | 0
| 20 |  |  | 0 |  | 0
| 25 |  |  | 0 |  | 0
| 30 |  |  | 0 |  | 0
| 35 | 34.9801 | 34.983 | 0.0028722 | 34.9763 | 0.0038084
| 40 |  |  | 0 |  | 0
| 45 | 44.989 | 44.9815 | 0.0074835 | 44.9909 | 0.001913
| hardstop high |  |  | 0 |  | 0

## Repeatabilty
Test |  | 6
Test description |  | Bidirectional repeatability
Approach 6 different positions form both directions, Repeat 6 times
Positions to appraoch |  | -5 | 0 | 10 | 20 | 30 | 40
Position | -5deg | Approach 6 different positions form both directions, Repeat 6 times -5: from below
Test | 1 | 2 | 3 | 4 | 5 | 6 |  | Tests | Comments
Laser scanner position | -5,0168 | -5,0185 | -5,0184 | -5,0197 | -5,0196 | -5,0196 |  | 6
Open loop counter | -61715.8347
Posital encoder | 802867 | 802871 | 802876 | 802887 | 802879 | 802884
Position | -5deg | Approach 6 different positions form both directions, Repeat 6 times -5: from above
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position | -5,0073 | -5,0064 | -5,0065 | -5,006 | -5,0062 | -5,0044
Open loop counter | -61715.8347
Posital encoder | 802567 | 802576 | 802571 | 802568 | 802564 | 802563
Position | 0deg | Approach 6 different positions form both directions, Repeat 6 times -5: from below
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position | 0.0173 | 0.0171 | 0.0173 | 0.0166 | 0.0164 | 0.017
Open loop counter | 0
Posital encoder | 690536 | 690543 | 690540 | 690551 | 690547 | 690555
Position | 0deg | Approach 6 different positions form both directions, Repeat 6 times -5: from above
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position | 0.0006 | 0.0015 | 0.0011 | 0.0001 | 0.0002 | 0.001
Open loop counter | 0
Posital encoder | 690152 | 690179 | 690168 | 690167 | 690163 | 690159
Position | 10deg | Approach 6 different positions form both directions, Repeat 6 times -5: from below
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position | 9.9904 | 9.99 | 9.9896 | 9.9987 | 9.989 | 9.9886
Open loop counter | 123431.6694
Posital encoder | 465664 | 465671 | 465676 | 465683 | 465684 | 465688
Position | 10deg | Approach 6 different positions form both directions, Repeat 6 times -5: from above
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position | 10.0046 | 10.0047 | 10.0051 | 10.0046 | 10.0048 | 10.0054
Open loop counter | 123431.6694
Posital encoder | 465307 | 465327 | 465320 | 465316 | 465312 | 465308
Position | 20deg | Approach 6 different positions form both directions, Repeat 6 times -5: from below
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position | 19.9969 | 19.9954 | 19.9948 | 19.9948 | 19.995 | 19.9949
Open loop counter | 246863.3387
Posital encoder | 241008 | 241019 | 241023 | 241027 | 241032 | 241035
Position | 20deg | Approach 6 different positions form both directions, Repeat 6 times -5: from above
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position | 20.016 | 20.011 | 20.0113 | 20.0117 | 20.0114 | 20.0115
Open loop counter | 246863.3387
Posital encoder |  |  | 240691 | 240687 | 240683 | 240679
Position | 30deg | Approach 6 different positions form both directions, Repeat 6 times -5: from below
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position | 29.9964 | 29.9968 | 29.9974 | 29.9969 | 29.9963 | 29.9964
Open loop counter | 370295.0081
Posital encoder | 16411 | 16403 | 16407 | 16411 | 16415 | 16416
Position | 30deg | Approach 6 different positions form both directions, Repeat 6 times -5: from above
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position | 30.0162 | 30.0168 | 30.0167 | 30.0167 | 30.017 | 30.0174
Open loop counter | 370295.0081
Posital encoder | 15960 | 15952 | 15947 | 15940 | 15936 | 15931
Position | 40deg | Approach 6 different positions form both directions, Repeat 6 times -5: from below
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position | 40.0132 | 40.0078 | 40.0078 | 40.0044 | 40.0069 | 40.0055
Open loop counter | 493726.6774
Posital encoder | 2147275180 | 2147275163 | 2147275167 | 2147275172 | 2147275176 | 2147275179
Position | 40deg | Approach 6 different positions form both directions, Repeat 6 times -5: from above
Test | 1 | 2 | 3 | 4 | 5 | 6
Laser scanner position | 40.0172 | 40.0221 | 40.0176 | 40.0224 | 40.0186 | 40.0249
Open loop counter | 493726.6774
Posital encoder | 2147274879 | 2147274875 | 2147274871 | 2147274867 | 2147274864 | 2147274860